## Slide 1
Status of the preamp development and cables for the silicon detectors
Monday 22 nd March 2021
Laurent Leterrier, Xavier Fléchard

## Slide 2: Test results of the first prototype : Amptek version
[WISArD, Visio conf meeting, Monday 22 nd March]
Resolution < 4,5 keVSi FWHM
at room temperature
Cd ≈ 100 pF + 10 cm cable
Optimal shaping time = 4 µs
- Measurements were performed with the faster system and pulse generator.
- Resolution for the rear side is about 20 keVsi FWHM in the same conditions except Cd ≈ 470 pF and optimal shaping time ≈ 7 µs.
- Power consumption ≈ 2,8 W / board
More in Marcin’s talk (test at AIFIRA with Silicon detectors & 700keV alpha beam)

## Slide 3: Second prototype board : CAEN Version
[WISArD, Visio conf meeting, Monday 22 nd March]
Advantage : Cost ≈ 200 € per unit ( AMPTEK : cost = 600 $ per unit)
Disadvantage : Size = 39 x 20 x 3 mm (AMPTEK : size = 17 x 10 x 3 mm)
Uncertainty : Operating temperature ?  to be checked !
Main features of the board:
- 6 channels with 6 charge sensitive amplifiers CAEN A1422H version F2
- For each channel :
- 400 mV/MeVSi  2.4 V for 6 MeVSi in CSA output
- CSA output connected directly to input FASTER
Choice of FASTER input range: 5V  Max 12.5 MeV, LSB 0.75 keV
2V  Max 5 MeV, LSB 0.30 keV
- Decay time ≈ 27 µs
- Power consumption ≈ 288 mW / channel  1.8 W / board
- NEC (Cd = 100 pF + Ccoax = 10 pF) ≈ 3 keV at room Temperature

## Slide 4
[WISArD, Visio conf meeting, Monday 22 nd March]
HV input (100 V max)
TEST input (50Ω)
Power ± 12 V
5 inputs for jonction side
6 outputs
1 input for ohmic side
CAEN preamplifiers
Test scheduled for the end of March
Second prototype board : CAEN Version
I/O connecteur

## Slide 5
[WISArD, Visio conf meeting, Monday 22 nd March]
Second prototype board : CAEN Version
Not final version,
depends on mecanical constraints !

## Slide 6
[WISArD, Visio conf meeting, Monday 22 nd March]
Cable and connection :
- Use of Fischer connector (105 serie) : 104 serie not compatible with the cable
Quotation : 3600 euros for Fischer connectors, 5100 euros for cables (full custom, 300 m)
The cost of the distribution boxes is not estimated
Delivery time : 14 weeks  decision to be made before the end of march !

## Slide 7
[WISArD, Visio conf meeting, Monday 22 nd March]
Connection proposal for the WISArD project :
Front panel for one distribution box
One distribution box for 4 detectors

## Slide 8
beta detector calibration
at low energy

## Slide 9: Need for accurate calibration at low energy
[WISArD, Visio conf meeting, Monday 22 March 2021]
Systematic error 10-3 (from 2018 test experiment)
What matters:
Knowledge of threshold
(direct effect)
Achieve very low threshold
(backscattering)
To reach .1% level:
Lower threshold down to 10 keV
(x2.5 on Bck scat error)
Threshold uncertainty of 1-1.5 keV
(x10 on Threshold error)
Need for a precise characterization of positron detector at very low energy

## Slide 10: Build low energy low intensity electron gun
[WISArD, Visio conf meeting, Monday 22 March 2021]
Simulations:
G4: no X rays on detector
SIMION: e- from anode only
Monoenergetic e- beam (0-30 keV)
~ 600 counts per second on detector
Test bench “CEBI”
(canon à electron basse intensité)

## Slide 11: First tests with PVT scintillator (EJ204)
[WISArD, Visio conf meeting, Monday 22 March 2021]

## Slide 12: Bkg subtraction & fit
[WISArD, Visio conf meeting, Monday 22 March 2021]
Fit with combination of 5 Gaussian functions
7 free parameters: x1,w1,A1,A2,A3,A4 & A5
With xn = n x x1, and wn = sqrt(n) x w1
For n e-, the response function is a convolution of n response functions obtained for 1 e-
G4 simulations: backscattering negligible
Mean Edep=29.35 keV

## Slide 13: Results for EJ204 scintillator
[WISArD, Visio conf meeting, Monday 22 March 2021]
30keV
25keV
20keV
15keV
10keV
Reproducibility limited by
BKG fluctuations
Gain fluctuations
2 sets of measurements
to be improved

## Slide 14: Extract Birks quenching coefficient
[WISArD, Visio conf meeting, Monday 22 March 2021]
Should allow to set a threshold around 10keV
with precision of at least 2-3 keV
L: light yield
S: scintillation efficiency

## Slide 15: Tests with YAP scintillator
[WISArD, Visio conf meeting, Monday 22 March 2021]
G4 simulations
Backscattering not negligible…
Better resolution but requires more complex fitting procedure

## Slide 16: Next
[WISArD, Visio conf meeting, Monday 22 March 2021]
Characterization of other PVT scintillators
(2018 experiment, Simon exp…)
Tests with SiPMs
Calibration of future beta detector for WISArD…
goal: threshold < 10 keV & ~1-1.5 keV precision (0.1% on abn)
Mid April?
Removable 241Am source
(ref for experiment @ ISOLDE)

## Slide 17: MCP detector
[WISArD, Visio conf meeting, Monday 22 March 2021]
Goal: monitoring implantation profile on catcher (resolution better than 300um)
Constraints: compact, works in strong Bfield
Resistive anode
2 MCPs
25mm
To be tested soon with mask & alpha source…